## Thera_Lase_Surgery_20220927
Date	Time	Omnia File Name	PC File Name	File	Step	Test Type	Pass/Fail	Test Result	Meter 1	Meter 2	Meter 3	Timer	Meter 4	Meter 5	Model Number	Serial Number	Probe	Measuring Device	Test Model #	Test Serial #	Calibration Due Date	Operator ID
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	1	Run Test	Pass	Pass	272.3 V	0.07 A	5 W	6.0 s	0.13 mA	0.262	OP - 72100	1880			Omnia 8106	9520337	11/07/2023	marilene	1
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	2	AC Withstand	Pass	Pass	3.63 kV	0.371 mA	0.063 mA	1.0 s			OP - 72100	1880			Omnia 8106	9520337	11/07/2023	marilene	2
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	3	AC Withstand	Pass	Pass	3.63 kV	0.114 mA	0.011 mA	1.0 s			OP - 72100	1880			Omnia 8106	9520337	11/07/2023	marilene	3
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	4	Line Leakage	Pass	Pass	272.6 V	20.8 uA		10.0 s			OP - 72100	1880	Probe-HI to Line	IEC60601, UL2601	Omnia 8106	9520337	11/07/2023	marilene	4
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	5	Line Leakage	Pass	Pass	272.5 V	20.9 uA		10.0 s			OP - 72100	1880	Probe-HI to Line	IEC60601, UL2601	Omnia 8106	9520337	11/07/2023	marilene	5
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	6	Line Leakage	Pass	Pass	272.6 V	13.8 uA		10.0 s			OP - 72100	1880	Probe-HI to Line	IEC60601, UL2601	Omnia 8106	9520337	11/07/2023	marilene	6
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	7	Line Leakage	Pass	Pass	272.7 V	13.7 uA		10.0 s			OP - 72100	1880	Probe-HI to Line	IEC60601, UL2601	Omnia 8106	9520337	11/07/2023	marilene	7
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	8	Line Leakage	Pass	Pass	272.6 V	13.4 uA		10.0 s			OP - 72100	1880	Probe-HI to Line	IEC60601, UL2601	Omnia 8106	9520337	11/07/2023	marilene	8
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	9	Line Leakage	Pass	Pass	272.6 V	13.3 uA		10.0 s			OP - 72100	1880	Probe-HI to Line	IEC60601, UL2601	Omnia 8106	9520337	11/07/2023	marilene	9
27/09/2022	07:52	TEMP/	TheraLaseSurgery_6-INSP-493_rev03.Omnia	1	10	Run Test	Pass	Pass	272.6 V	0.07 A	5 W	6.0 s	0.13 mA	0.262	OP - 72100	1880			Omnia 8106	9520337	11/07/2023	marilene	10